Supplements

An extremely sensitive amylase activity assay and its application for the determination of the residual amylase activity in bread

Katrin Reichenberger, Sabine Lutz-Wahl, Lucas Kettner and Lutz Fischer*

University of Hohenheim, Institute of Food Science and Biotechnology,

Department of Biotechnology and Enzyme Science, Garbenstr. 25, 70599 Stuttgart, Germany

*Corresponding author: Lutz Fischer

E-mail address: Lutz.Fischer@uni-hohenheim.de

Tel.: +49 711 459-22311

Keywords: amylase assay, α-amylase, thermal stability, residual activity, bread baking, enzyme declaration, food enzyme

Table S1: Determination of different endo-amylase activities using different blanks (B1 and B2)

| Submitted endo-amylase activity [pkat mL-1] | Activity calculated with Blank 1*1 [pkat mL-1] | Activity calculated with Blank 2*2 [pkat mL-1] | Percentage activity with Blank 1 [%] | Percentage activity with Blank 2 [%] | Assay run time [h] |
| --- | --- | --- | --- | --- | --- |
| 60 | 60.91 ± 1.54 | 58.11 | 101.5 | 96.5 | 1.25 |
| 15 | 26.52 ± 0.98 | 15.35 | 176.8 | 102.3 | 2 |
| 10 | 18.20 ± 0.886 | 9.20 | 182.0 | 92.0 | 3 |
| 2 | 3.36 ± 0.074 | 1.94 | 168.0 | 97.0 | 17 |
| 1 | 3.73 ± 0.0303 | 0.93 | 373.0 | 93.0 | 24 |
| 0.5 | <LOD | <LOD |  |  | 48 |

*1 incubation of the Ceralpha reagent for the respective assay run time, addition of the stopping reagent followed by the addition of amylase solution

*2 incubation of the Ceralpha reagent together with heat inactivated amylase solution or buffer for the respective assay run time, addition of the stopping reagent

Figure S1: First calibration curve of the new Da67 assay using maltose (concentrations of assay components not yet optimized: 50 μM Da67, 681 nkat mL-1 peroxidase, 150 nkat mL-1 glucose oxidase and 1 % (w/v) starch)

Figure S2: Calibration curve of the new Da67 assay using maltose with optimized concentration of assay components (50 μM Da67, 681 U mL-1 peroxidase, 93 U mL-1 glucose oxidase and 0.01 % (w/v) starch)

Figure S3: Absorbance values of one exemplary sample (50 ng mL-1 of the α-amylase from B. subtilis) and one assay blank using a long assay run time with all assay components together (Da67, peroxidase, glucose oxidase, starch and amylase)